## Instructions
Excel book for the calculation of pH measured by spectrophotometry and uncertainties by the Kragten method
Please cite as: Cartas H., Sanchez-Cabeza J.A., Martínez, R., Gómez M., Perez, E., Herrera J., Ruiz G., & Bernal C. (2021) Excel book for the calculation of pH measured by spectrophotometry and uncertainties by the Kragten method. Red de Investigación de Estresores Marinos - Costeros en Latinoamérica y El Caribe – REMARCO. Cienfuegos, Cuba. https://remarco.org/manual-ao/
This follows the procedures published by REMARCO (https://remarco.org/manual-ao/)
Please read the manual
Written by Héctor Cartas, CEAC, 2021-04-08
Adapted by Prof. Joan-Albert Sanchez-Cabeza, 2021-04-09
Sheet Input: sample data and measurements. This is the only sheet where the user enters information
Sheet Results: final results are presented
Sheet Kragten-pH (occult): the Kragten method is used for uncertainty calculations
To avoid accidental changes of calculation cells, only the worksheet Input is unlocked
To unlock, the password is remarco (lowercase)
Please, send your comments to Hector Cartas Aguila <hector@ceac.cu>
This work was supported by the International Atomic Energy Agency - IAEA, through the project RLA/7/025 entitled "Strengthening Capacities in Marine and Coastal Environments Using Nuclear and Isotopic Techniques".

## Input
| Salinity (psu) | Temperature (°C) | Volumes: |  | Absorbance (1) | Absorbance (1) | Absorbance (1) | Absorbance (2) | Absorbance (2) | Absorbance (2) | Resolution: |  | Salinity (PSU) | Temperature (°C) | Volume 1 (mL) | Volume 2 (mL) | Absorbance (1) | Absorbance (1) | Absorbance (1) | Absorbance (2) | Absorbance (2) | Absorbance (2) | pH | Repetibility:
Code | Sample | Analysis |  |  | 434 nm | 578 nm | 730 nm | 434 nm | 578 nm | 730 nm |  |  | Repetitions | Repetitions | Repetitions | Repetitions | 434 nm repetitions | 578 nm repetitions | 730 nm repetitions | 434 nm repetitions | 578 nm repetitions | 730 nm repetitions | Repetitions | Salinity (PSU)
ph.ceac.1 | 36.424 | 25 | Volume 1 (mL) | 0.05 | 0.23342 | 0.52847 | 0.0061947 | 0.44154 | 0.98962 | 0.0074553 | Salinity (PSU) | 0.01 |  |  | 0.001304 | 0.001304 | [7 empty] | Temperature (°C)
| 36.412 | 25 | Volume 2 (mL) | 0.05 | 0.23312 | 0.52828 | 0.005692 | 0.44182 | 0.99099 | 0.0073837 | Temperature (°C) | 0.1 |  |  | 0.0018056 | 0.0018056 | [7 empty] | Volume 1 (mL) | 0.0003
| 36.409 | 24.9 |  |  | 0.23249 | 0.52768 | 0.004988 | 0.44199 | 0.9895 | 0.0070757 | Volume 1 (mL) | 0.0001 |  |  | 0.0023071 | 0.0023071 | [7 empty] | Volume 2 (mL) | 0.0003
| 36.41 | 24.9 |  |  | 0.23185 | 0.52738 | 0.0050253 | 0.44194 | 0.98969 | 0.0068987 | Volume 2 (mL) | 0.0001 |  |  | 0.0015046 | 0.0015046 | [7 empty] | Absorbance (1, 434)
| 36.409 | 25 |  |  | 0.23157 | 0.52676 | 0.0050183 | 0.44195 | 0.99022 | 0.0081397 | Absorbance | 1e-06 |  |  | 0.0017053 | 0.0017053 | [7 empty] | Absorbance (1, 578)
| 36.413 | 25.1 |  |  | 0.231 | 0.52507 | 0.005076 | 0.44252 | 0.98967 | 0.0079413 |  |  |  |  | 0.0019059 | 0.0019059 | [7 empty] | Absorbance (1, 730)
| 36.41 | 25 |  |  | 0.22951 | 0.52335 | 0.0052883 | 0.44231 | 0.98881 | 0.007325 |  |  |  |  | 0.001605 | 0.001605 | [7 empty] | Absorbance (2, 434)
| 36.406 | 25 |  |  | 0.22897 | 0.52153 | 0.0053047 | 0.44175 | 0.9884 | 0.0063463 |  |  |  |  | 0.0018056 | 0.0018056 | [7 empty] | Absorbance (2, 578)
| 36.414 | 25 |  |  | 0.22789 | 0.51947 | 0.0054733 | 0.44053 | 0.98819 | 0.0058793 |  |  |  |  | 0.0017053 | 0.0017053 | [7 empty] | Absorbance (2, 730)
| 36.412 | 25 |  |  | 0.22707 | 0.51778 | 0.0052677 | 0.43992 | 0.98738 | 0.0059237 |  |  |  |  | 0.0019059 | 0.0019059 | [7 empty] | pH | 0.0026

## Results
SUMMARY | [13 empty] | UNCERTAINTY BUDGET (PLOT)
pH = | 8.0783
u(pH) = | 0.0044052
ue(pH) = | 0.0088103
Quality index = | 96.6667 | %
UNCERTAINTIES | Salinity (PSU) | Temperature (°C) | Volume 1 (mL) | Volume 2 (mL) | Absorbance (1) | Absorbance (1) | Absorbance (1) | Absorbance (2) | Absorbance (2) | Absorbance (2) | pH
[5 empty] | 434 nm | 578 nm | 730 nm | 434 nm | 578 nm | 730 nm
Numbers:
n sample | 10 | 10 | 0 | 0 | 10 | 10 | 10 | 10 | 10 | 10 | 0
n repetitions | 0 | 0 | 10 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0
n repetibility | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1
Repetibility:
Sample | 0.0013974 | 0.016387 |  |  | 0.00063921 | 0.0011162 | 0.00010874 | 0.00023191 | 0.00030553 | 0.00022545
Repetitions | 0 | 0 | 5.4815e-05 | 5.4815e-05 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Repetibility | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0026
Uncertainties:
Mean | 36.4119 | 24.99 | 0.05 | 0.05 | 0.23069 | 0.52458 | 0.0053328 | 0.44163 | 0.98925 | 0.0070369
u(repetibility) | 0.0013974 | 0.016387 | 5.4815e-05 | 5.4815e-05 | 0.00063921 | 0.0011162 | 0.00010874 | 0.00023191 | 0.00030553 | 0.00022545 | 0.0026
u(resolution) | 0.0028868 | 0.028868 | 2.8868e-05 | 2.8868e-05 | 2.8868e-07 | 2.8868e-07 | 2.8868e-07 | 2.8868e-07 | 2.8868e-07 | 2.8868e-07
u(combined) | 0.0032072 | 0.033194 | 6.1951e-05 | 6.1951e-05 | 0.00063921 | 0.0011162 | 0.00010874 | 0.00023191 | 0.00030553 | 0.00022545 | 0.0026
Quality index:
Maximum | 3 | 3 | 2 | 2 | 3 | 3 | 3 | 3 | 3 | 3 | 2 | 30 |  | Maximum value
Position | 1 | 1 | 2 | 2 | 1 | 1 | 1 | 1 | 1 | 1 | 3
Value | 3 | 3 | 2 | 2 | 3 | 3 | 3 | 3 | 3 | 3 | 1 | 29 |  | Sum
[12 empty] | 96.6667 | % | Quality index

## Kragten-pH (hidden)
|  | u[S](psu) | u(T)(K) | u(A730, 1 mCP) | u(A578, 1 mCP) | u(A434, 1 mCP) | u(A730, 2 mCP) | u(A578, 2 mCP) | u(A434, 2 mCP) | u(V1)(mL) | u(V2)(mL)
|  | 0.0032072 | 0.033194 | 0.00010874 | 0.0011162 | 0.00063921 | 0.00022545 | 0.00030553 | 0.00023191 | 6.1951e-05 | 8.7613e-05
S(psu) | 36.4119 | 36.4151 | 36.4119 | 36.4119 | 36.4119 | 36.4119 | 36.4119 | 36.4119 | 36.4119 | 36.4119 | 36.4119
T(K) | 298.14 | 298.14 | 298.1732 | 298.14 | 298.14 | 298.14 | 298.14 | 298.14 | 298.14 | 298.14 | 298.14
A730, 1 mCP | 0.0053328 | 0.0053328 | 0.0053328 | 0.0054416 | 0.0053328 | 0.0053328 | 0.0053328 | 0.0053328 | 0.0053328 | 0.0053328 | 0.0053328
A578, 1 mCP | 0.52458 | 0.52458 | 0.52458 | 0.52458 | 0.52569 | 0.52458 | 0.52458 | 0.52458 | 0.52458 | 0.52458 | 0.52458
A434, 1 mCP | 0.23069 | 0.23069 | 0.23069 | 0.23069 | 0.23069 | 0.23133 | 0.23069 | 0.23069 | 0.23069 | 0.23069 | 0.23069
A730, 2 mCP | 0.0070369 | 0.0070369 | 0.0070369 | 0.0070369 | 0.0070369 | 0.0070369 | 0.0072623 | 0.0070369 | 0.0070369 | 0.0070369 | 0.0070369
A578, 2 mCP | 0.98925 | 0.98925 | 0.98925 | 0.98925 | 0.98925 | 0.98925 | 0.98925 | 0.98955 | 0.98925 | 0.98925 | 0.98925
A434, 2 mCP | 0.44163 | 0.44163 | 0.44163 | 0.44163 | 0.44163 | 0.44163 | 0.44163 | 0.44163 | 0.44186 | 0.44163 | 0.44163
V1(mL) | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.050062 | 0.05
V2(mL) | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.10009
pH | 8.0783 | 8.0783 | 8.0779 | 8.0786 | 8.0804 | 8.0755 | 8.0782 | 8.0782 | 8.0786 | 8.0783 | 8.0783
|  | Salinity | Temperature | A730,1 | A578,1 | A434,1 | A730,2 | A578,2 | A434,2 | V1 | V2 | pH
| Uncertainty: | u[pH(S)] | u[pH(T)] | u[pH(A730, 1)] | u[pH(A578, 1)] | u[pH(A434, 1)] | u[pH(A730, 2)] | u[pH(A578, 2)] | u[pH(A434, 2)] | u[pH(V1)] | u[pH(V2)] | u[pH(Rep)]
|  | -3.0661e-06 | -0.00041853 | 0.00026931 | 0.0021155 | -0.0027951 | -0.00013994 | -0.00015043 | 0.00025784 | 2.3362e-05 | -1.6471e-05 | 0.0026
| Variance: | u[pH(S)]^2 | u[pH(T)]^2 | u[pH(R1)]^2 | u[pH(A578, 1mCP)]^2 | u[pH(A434, 1mCP)]^2 | u[pH(A730, 2mCP)]^2 | u[pH(A578, 2mCP)]^2 | u[pH(A434, 2mCP)]^2 | u[pH(V1)]^2 | u[pH(V2)]^2 | u[pH(Repetibilidad)]^2 | Total variance
|  | 9.4009e-12 | 1.7517e-07 | 7.2529e-08 | 4.4755e-06 | 7.8127e-06 | 1.9583e-08 | 2.2629e-08 | 6.648e-08 | 5.4578e-10 | 2.7128e-10 | 6.76e-06 | 1.9405e-05
| Relative variance: | 4.8445e-05 | 0.90269 | 0.37376 | 23.0632 | 40.2603 | 0.10092 | 0.11661 | 0.34259 | 0.0028125 | 0.001398 | 34.8357
u(pH) | 0.0044052
ur(pH)(%) | 0.054531